GEEG-ZUNA: Evaluation of a Large-Scale EEG Foundation Model for Channel Reconstruction

Project Overview

This project evaluates ZUNA (Zyphra, arXiv 2602.18478) — a 380-million parameter masked diffusion autoencoder trained on large-scale EEG corpora — for the task of reconstructing corrupted or missing electrode channels from resting-state human EEG. The primary research question is whether ZUNA's learned spatiotemporal priors outperform classical spherical spline interpolation (the current clinical and research standard) when a subset of channels are absent.

Model Under Test

ZUNA encodes EEG using 4D Rotary Positional Encoding over electrode (x, y, z) position and time. It tokenizes EEG into coarse spatial–temporal blocks (128 fine timepoints per coarse step; 10 coarse steps per 5 s epoch at 256 Hz) and reconstructs masked channels via diffusion sampling (50 steps used here). The model was trained on average-referenced 256 Hz data. Input normalization requires z-scoring followed by division by 10 (data_norm=10.0) to bring the signal to the model's expected std ≈ 0.1.

Data

- Subject: G001, single resting-state session (Day1Rest1)

- Format: Neuroscan .cnt file

- Channels: 62-channel full-cap EEG (10-20 system, standard_1005 montage)

- Sample rate: 256 Hz (resampled from acquisition rate)

- Epochs: 16 × 5.0 s epochs, marker-locked to event codes broadcast every 0.5 s during confirmed resting state to avoid artifact-contaminated periods

- Preprocessing: 1–40 Hz bandpass filter, resampled to 256 Hz, edge artifact crop, average reference applied (required to match ZUNA training distribution)

Pipeline Architecture

main_pipeline.py — Four-stage orchestrator:

| Stage | Description |
| --- | --- |
| 0 | Load CNT → epoch → export ground truth tensors |
| 1 | Package .pt tensors for ZUNA inference |
| 2 | ZUNA inference (inference(), 50 diffusion steps, data_norm=10.0) |
| 3 | Z-score reversal, hard inpainting (preserve observed channels exactly), export .npy |
| 4 | Spectral analysis — IAF (center-of-gravity, posterior channels, 0.125 Hz resolution) and FAA (F3/F4 log power ratio) |

Baseline: MNE interpolate_bads() spherical spline interpolation, the accepted standard in BCI and clinical EEG.

Hard inpainting is applied post-inference: ZUNA reconstructs all channels, but the observed (non-dropped) channels are overwritten with the original ground truth signal, ensuring only the missing channels are evaluated.

Evaluation Design

Primary Test: N-Channel Dropout Sweep (test_n_channel_dropout.py)

Channels are randomly dropped (N = 2, 4, 8) with 8 independent trials per N value (different random seeds per trial), producing 24 total inference runs. Each trial compares ZUNA reconstruction against spline interpolation on the dropped channels only.

Metrics per trial:

- Temporal correlation (Pearson r, per-epoch mean across dropped channels)

- RMSE (µV)

- Signal-to-Distortion Ratio (SDR, dB)

- Spectral correlation (Pearson r on PSD)

- Band power error (mean absolute log-ratio, five bands: δ, θ, α, β, γ)

Results are aggregated across 8 trials per N and visualized with mean ± SD error bars in SWEEP_SUMMARY.png.

Secondary Analysis: Full-Channel Comparison (compare_outputs.py)

13 figure panels covering the original 19→64 channel reconstruction test:

- Temporal and spectral correlation per channel

- RMSE and relative RMSE

- Band power error (all five bands)

- SDR

- Spatial correlation

- Per-channel time series panels (all 62 channels, multiple epochs)

- Functional connectivity: PLV and wPLI in alpha and beta bands

- ICA topographic maps

- Source-space ROI reconstruction error

Key Results (N-Channel Dropout Sweep)

Aggregated across 8 trials per condition (single subject, single session):

| Metric | N | Spline | ZUNA |
| --- | --- | --- | --- |
| Temporal r | 2 | 0.674 | 0.759 |
| Temporal r | 4 | 0.653 | 0.732 |
| Temporal r | 8 | 0.685 | 0.733 |
| RMSE (µV) | 2 | 3.13 | 2.69 |
| RMSE (µV) | 4 | 4.16 | 3.17 |
| RMSE (µV) | 8 | 4.29 | 4.23 |
| Beta BP error | 2 | 1.419 | 0.124 |
| Gamma BP error | 2 | 1.674 | 0.386 |
| SDR (dB) | 2 | 3.40 | 4.04 |
| SDR (dB) | 4 | 1.79 | 3.27 |
| SDR (dB) | 8 | 4.29 | 3.42 |

ZUNA wins consistently on temporal fidelity, RMSE, and high-frequency band power preservation across all three N values. Spline SDR is marginally better at N=8, likely because spline leverages all 54 surrounding channels as constraints while ZUNA's long-range prior degrades slightly with heavier masking.

Alpha Metrics (IAF / FAA — 19-channel original test)

| Condition | IAF (Hz) | FAA |
| --- | --- | --- |
| Truth | 10.379 | — |
| Spline | 10.369 | — |
| ZUNA | 10.363 | — |

IAF computed via center-of-gravity on posterior channels (O1, O2, Oz, P3, P4, Pz, P7, P8, PO3, PO4, POz) at 0.125 Hz resolution. Both methods preserve IAF faithfully; differences are sub-bin and clinically negligible.

Critical Bugs Identified and Corrected

| Bug | Impact | Fix |
| --- | --- | --- |
| data_norm not passed | ZUNA received std≈1.0 instead of expected std≈0.1 — 10× scale error on all inference | Added data_norm=10.0 explicitly to inference() call |
| No average reference | Input data in mastoid/nose reference; ZUNA trained on avg ref | raw.set_eeg_reference('average') added in load_data.py |
| sfreq=250 default in get_alpha_metrics | Frequency axis mis-scaled for PSD and IAF | Corrected default to 256 Hz |
| IAF from all channels, argmax peak-picking | Frontal/temporal alpha diluted IAF; 0.5 Hz grid too coarse | Posterior-only center-of-gravity at 0.125 Hz resolution |
| Token count comment incorrect | Documentation only; execution unaffected | Corrected to 10 coarse steps (1280/128) |

Software Environment

- Python 3.9, PyTorch, MNE-Python 1.x

- ZUNA inference via zuna.inference() (Zyphra open-source release)

- GPU: single NVIDIA consumer GPU (Windows 10, ReFS volume)

- HuggingFace model cache: D:\GEEG-ZUNA\HF_cache

Limitations

- Single subject, single session — results cannot be generalized; cross-subject replication is required.

- Resting state only — ZUNA's advantage may differ for event-related or task-evoked paradigms where transient non-stationary signals dominate.

- N=8 SDR reversal — at higher dropout rates, the advantage margin narrows. The N=16 and N=32 regimes remain untested.

- Windows/ReFS instability — repeated GPU crashes from ZUNA subprocess accumulated kernel-object corruption on the ReFS filesystem, requiring reboots between run batches. This is an infrastructure limitation, not a model limitation.

Outputs

| File/Directory | Contents |
| --- | --- |
| results_n_channel_dropout/SWEEP_SUMMARY.png | Error-bar summary across N=2,4,8 |
| results_n_channel_dropout/N0{2,4,8}_drop/trial_{01-08}/ | Per-trial time series, PSD, epoch correlation figures |
| comparison_results/ | 13 full-pipeline comparison figures (19→64 channel test) |
| alpha_comparison.png | PSD overlay with IAF/FAA annotation |
| pipeline_audit.log | Full run log with health checks |
| G001_Day1Rest1_X_zuna_test.npy | Final ZUNA reconstruction array (16 × 62 × 1280) |